## Calibration
Calibration Log
Calibration shall be made every 12 months or after repair for each CRM by either, the monitor manufacturer or a national radon proficiency program approved calibration laboratory approved by the device manufacturer. Any monitor that does not have a calibration certification, dated within 12 months, must be removed from service. In addition to calibration, an annual background check is performed by purging with clean aged air or nitrogen. The manufacturer or calibration laboratory completes this process at the time of calibration.
Monitor Serial Number | Calibration Date 2021 | Calibration Date 2022 | Calibration Date 2023 | Calibration Date 2024 | Calibration Date 2025 | Calibration Date 2026
ABC123 | 2018-01-06
ABD321 | 2018-04-15
DCB456 | 2018-07-24
CDA233 | 2018-11-10

## Dups <2
Difference Log for Duplicates Averaging Less Than 2.0 pCi/L
<Do not enter data in the columns labeled Difference since it has a formula to complete this calculation. In this worksheet, you can extend the formulas in the Difference column by highlighting the cells in the upper portions of the column and dragging the bottom right corner of the column downward. In this example, we have entered data in the first or top five cells of the radon concentrations columns as illustrations or examples.>
Start Date | End Date | Higher Test Result - Device Serial Number | Lower Test Result - Device Serial Number | Higher Test Result - Radon Concentration (pCi/L) | Lower Test Result - Radon Concentration (pCi/L) | <Do Not Enter Data in this Columns. The information is calculated by the formula already in the template> Difference | (Warning limit is a difference of 1 pCi/L when Investigation is needed)
Notes
2017-01-01 | 2017-01-03 | 123456789 | 34567891011 | 1.8 | 1.4 | 0.4
2017-01-02 | 2017-01-04 | 987654321 | 683455320 | 1.5 | 0.8 | 0.7
2017-01-03 | 2017-01-05 | 2345678910 | 89452757 | 1.9 | 1.1 | 0.8
2017-01-04 | 2017-01-06 | 1098765432 | 7753066382 | 1.4 | 0.3 | 1.1 | <Investigation needed>
2017-01-06 | 2017-01-07 | 1098765432 | 7753066382 | 0.9 | 0.5 | 0.4
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Dups 2.0-3.9
Relative Percent Difference Log and Chart for Duplicates Averaging 2 to 3.9 pCi/L
<Do not enter data in the columns labeled Difference, Average, or RPD since they have formulas to complete these calculations. In this spreadsheet, you can extend the formulas in the Difference, Average, and RPD columns by highlighting the cells in the upper portions of each column and dragging the bottom right corner of the column downward. In this example, we have entered data in the first or top five cells of the radon concentrations columns to illustrate how the table is tied to the linked chart. When you are ready to enter your own data, you will want to replace the data in these cells.>
Start Date | End Date | Higher Test Result - Device Serial Number | Lower Test Result - Device Serial Number | Higher Test Result - Radon Concentration (pCi/L) | Lower Test Result - Radon Concentration (pCi/L) | <Do Not Enter Data in these Columns. The information is calculated by the formulas already in the template> |  |  | >67% Control Limit
50 to 67% Warning Level
26 to 49% In Control
0 to 25% Expected Precision
[6 empty] | Difference | Average | RPD | Notes:
2017-01-01 | 2017-01-03 | 123456789 | 34567891011 | 4 | 3.8 | 0.2 | 3.9 | 5.128%
2017-01-02 | 2017-01-04 | 987654321 | 683455320 | 3.9 | 3.3 | 0.6 | 3.6 | 16.67%
2017-01-03 | 2017-01-05 | 2345678910 | 89452757 | 2.2 | 1.8 | 0.4 | 2 | 20%
2017-01-04 | 2017-01-06 | 1098765432 | 7753066382 | 3.8 | 2 | 1.8 | 2.9 | 62.07%
2017-01-07 | 2017-01-07 | 1098765432 | 29648656435 | 4 | 1.9 | 2.1 | 2.95 | 71.19% | 1098765432 sent in for calibration
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!

## Dups 4+
Relative Percent Difference Log and Chart for Duplicates Averaging 4 pCi/L or More
<Do not enter data in the columns labeled Difference, Average, or RPD since they have formulas to complete these calculations. In this spreadsheet, you can extend the formulas in the Difference, Average, and RPD columns by highlighting the cells in the upper portions of each column and dragging the bottom right corner of the column downward. In this example, we have entered data in the first or top five cells of the radon concentrations columns to illustrate how the table is tied to the linked chart. When you are ready to enter your own data, you will want to replace the data in these cells.>
Start Date | End Date | Higher Test Result - Device Serial Number | Lower Test Result - Device Serial Number | Higher Test Result - Radon Concentration (pCi/L) | Lower Test Result - Radon Concentration (pCi/L) | <Do Not Enter Data in these Columns. The information is calculated by the formulas already in the template> |  |  | >36% Control Limit
28 to 36% Warning Level
15 to 27% In Control
0 to 14% Expected Precision
[6 empty] | Difference | Average | RPD | Notes:
2017-01-01 | 2017-01-03 | 123456789 | 34567891011 | 22.6 | 22 | 0.6 | 22.3 | 2.691%
2017-01-04 | 2017-01-06 | 123456789 | 34567891011 | 6.8 | 6.6 | 0.2 | 6.7 | 2.985%
2017-01-03 | 2017-01-05 | 2345678910 | 89452757 | 4.9 | 4.4 | 0.5 | 4.65 | 10.75%
2017-01-04 | 2017-01-06 | 1098765432 | 7753066382 | 7.2 | 5.2 | 2 | 6.2 | 32.26% | Conducted another duplicate with CRM and was in the expected precision range
2017-01-05 | 2017-01-07 | 34567891011 | 29648656435 | 7.1 | 6.8 | 0.3 | 6.95 | 4.317%
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!
[6 empty] | 0 | 0 | #DIV/0!

## Blanks
Field Blank Control Log and Chart for Passive, Time Integrating Devices
<Note, in this example, the device had a Lower Limit of Detection of 0.3 pCi/L. Obtain the LLD from the device manufacturer for the device you use and insert the values in this table.>
Start Date | End Date | Device Identification Number | Radon Concentration (pCi/L) |  | Investigation needed for anything above lower limit of detection
|  |  |  | Location (e.g. field, storage, lab-transit) | Notes
2017-01-01 | 2017-01-03 | 123456789 | 0.4 |  | Investigation needed
2017-01-02 | 2017-01-04 | 987654321 | <0.3
2017-01-03 | 2017-01-05 | 2345678910 | 0.8 |  | Contact lab. Test storage area and elevated radon found. Kits were discarded.
2017-01-04 | 2017-01-06 | 1098765432 | 0.5 |  | Investigation needed
2017-01-05 | 2017-01-07 | 34567891011 | <0.3 |  | Example

## Spikes
Field Spike Control Log and Chart for Passive Time Integrating Devices
<Do not enter data in the column labeled Relative Percent Error since it has formulas to complete these calculations. In this spreadsheet, you can extend the formula in the Relative Percent Error column by highlighting the cells in the upper portions of each column and dragging the bottom right corner of the column downward. In this example, we have entered data in the first or top five cells of the radon concentrations columns to illustrate how the table is tied to the linked chart. When you are ready to enter your own data, you will want to replace the data in these cells. Relative Percent Error (RPE) is calculated by subtracting the difference between the spiking chamber’s value and that obtained from the analytical laboratory with the difference divided by the spiking chamber’s value. >
Start Date | End Date | Device Identification Number | Spike Number | Measured Value (MV) (pCi/L) (Test Result from the Lab) | Reference Value (RV) (pCi/L) (Stated Radon Level from the Spiking Chamber) | <Do Not Enter Data in these Columns. The information is calculated by the formulas already in the template> |  | 31% Control Limit
21% to 30% Warning Level
11 to 20% In Control
-10 to +10% Expected Precision
-11 to -20% In Control
-21% to -30% Warning Level
-31% Control Limit
[6 empty] | Relative Percent Error (RPE) | Notes
2017-01-01 | 2017-01-03 | ABC123 | 0001 | 9.1 | 8.9 | 2.247%
2017-01-02 | 2017-01-04 | ABD321 | 0002 | 11.4 | 14.6 | -21.92%
2017-01-03 | 2017-01-05 | DCB456 | 0003 | 12.1 | 8.6 | 40.7% |  | Contacted lab. Lab said to return kits and sent a new batch. Spikes on the new batch were in control.
2017-01-04 | 2017-01-06 | CDA233 | 0004 | 6.4 | 6.5 | -1.538%
2017-01-05 | 2017-01-07 | EDA541 | 0005 | 10.6 | 9.3 | 13.98%
|  |  | 0006 | 10.6 | 9.4 | 12.77%
|  |  | 0007 |  |  | #DIV/0!
|  |  | 0008 |  |  | #DIV/0!
|  |  | 0009 |  |  | #DIV/0!
|  |  | 0010 |  |  | #DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!
#DIV/0!